## Slide 1
SLAC NATIONAL ACCELERATOR LABORATORY · SSRL BEAMLINE 1-5
Autonomous SWAXS Nanoparticle
Synthesis Platform
Closed-loop flow synthesis with in-situ SAXS/WAXS — beamtime report
Prepared for: PI & synthesis / beamline team
Scope: full pipeline + beamtime work, bench validation, and next steps
Notes: Title. We built and bench-validated an autonomous flow-synthesis + in-situ SAXS platform on BL1-5; this deck walks the team through the whole pipeline and what we accomplished this beamtime.

## Slide 2
EXECUTIVE SUMMARY
What we built and validated
9
web apps
one hub, closed-loop pipeline
135
automated tests
green (golden pipeline + safety)
3 / 4
beamline links live
reads · temp · shutter · collect
10+
docs & audits
setup, safety, readiness, design
The one-paragraph version
We built an autonomous platform that runs a flow-synthesis reactor, triggers in-situ SAXS/WAXS collection through SPEC, and feeds reduced → averaged → background-subtracted profiles to a nanoparticle-fit + Bayesian optimizer that proposes the next synthesis conditions — closing the loop.
This beamtime we integrated and bench-validated the beamline control (live monitors, temperature, shutter, 2D collection), added EPICS live reads, per-pump flow calibration, and hardware workarounds, hardened safety, and produced an end-to-end readiness audit. The autonomous data path is wired; the run-time collection trigger is the remaining item under debug.
Notes: Lead with outcomes: a working, tested, documented closed-loop platform; beamline control bench-validated; one open item.

## Slide 3
ARCHITECTURE
Hub-and-spoke: nine focused apps
All science lives in a shared src/ package; each app is a thin Flask shell. One hub launches & monitors them and shares the project folder.
Hub :5000
Launcher, project folder, event bus
Reduction :5001
Raw 2D → 1D I(q), corrections (PyFAI)
Viewer :5002
View & average repeated scans
Background :5003
Subtract buffer; pair by recipe_id
Analysis :5004
Guinier / Porod / Kratky / peaks
Assistant :5005
AI Q&A, plots, quality hints
Quality :5006
Good / NeedsReview grading
Reactor :5007
Pumps + beamline control, autopilot
Analyzer :5008
NP fit + Bayesian optimizer (loop brain)
Notes: Nine apps, hub-and-spoke. Design rule: all logic in src/, apps are thin shells.

## Slide 4
THE CLOSED LOOP
How an autonomous cycle flows
Reactor
arm → run → flush
SPEC 2D collect
sample + background
Reduce
2D → 1D I(q)
Average
repeat scans
Background
subtract (pair by id)
Analyzer
size / PDI / phase
Optimizer
next conditions
recipe_id ties every shot to its data all the way to the next proposed recipe → full traceability & sample↔background pairing.
Notes: Walk the loop; the teal box (optimizer) writes a condition file the reactor watches — the closing link. recipe_id is the traceability handle.

## Slide 5
HARDWARE & CONTROL
Topology: control PC ↔ bServer ↔ SPEC
Control PC
Hub + reactor app; USB to 5 Mitos pumps; HTTP to bServer
bServer (Windows)
pySSRL-bServer, 127.0.0.1:18085 — relays to SPEC
SPEC host (Linux)
Shutter, counters, SAXS+WAXS detectors; writes /msd_data
Fluidics & data path
5 Mitos P-pumps + LG16 sensors → mixer → heated reactor coil → flow cell (in beam).
Temperature commanded/read through SPEC (csettemp / CTEMP).
SPEC saves .raw to /msd_data/.../2D/{SAXS,WAXS}.
Same folder is the PC's X:\ mount — pipeline reads it back (Win↔Linux path map).
Notes: Control and data live on different hosts/paths. The X:↔/msd_data mapping is why data_dir translation matters.

## Slide 6
REACTOR CONTROL
State machine + arming
idle
arming
running
flushing
ready
E-stop from any state → pumps only
Arming modes
Temperature (reach & hold T) or Timed wait. (Ramp mode removed per request.)
Cooldown
csettemp → room temp the moment synthesis ends (run→flush).
Flush
Dedicated ode_flush, or a reagent pump (ode_dilution) — switchable in-app.
Live inputs
App values (duration, flush, exposure, keywords, data_dir) win over recipe & config, and stick for the session.
Notes: idle→arming→running→flushing→ready, E-stop anywhere. App inputs are authoritative & session-sticky; ramp arming removed.

## Slide 7
BEAMLINE INTEGRATION
SPEC control — safe by design
E-stop = pumps only
Never sends anything to the beamline; an in-progress X-ray collection is never interrupted.
Hard SPEC guard
One lock serializes SPEC; collect holds it the whole acquisition; live reads skip, never stall.
Auto remote control
Acquires SPEC control before any command; releases it on app exit so staff regain control.
No config writes
Only csettemp / ct / sopen-sclose / the collect macro — same commands as the group's own scripts.
Skips absent pumps
Disconnected/unused pumps (e.g. no-air ode_flush) aren't opened or polled — can't trip E-stop.
Notes: Nothing we do can crash SPEC or a running collection; control handed back cleanly; no persistent config changes (audited).

## Slide 8
DATA COLLECTION
Triggering SPEC in step with the run
How a shot is fired
Sample shot: spec_lead seconds before the run ends.
Background: spec_lead before the flush ends (clean solvent in cell).
Filenames <recipe_id>_sample / _bkg → pairing + traceability.
'Collect now' button for a manual, guarded acquisition.
Macro → SPEC (commands mode)
newfile → pd save → sopen →
loopscan <frames> <exposure> →
sclose → pd nosave
Streamed line-by-line via the bServer (no shared file needed).
Flat, values-inlined macro — avoids SPEC variable-vs-interactive pitfalls.
Defaults this beamtime: 10 frames × 10 s.
Notes: Commands mode + flat macro avoids the SPEC script-vs-interactive parsing issue and needs no shared-mount file.

## Slide 9
LIVE MONITORING
Temperature · I0 · beamstop — via EPICS
EPICS caget (recommended)
Reads PVs directly — independent of SPEC.
No remote control, no 'ct' pulses on the shutter.
Keeps updating DURING a data collection.
CA address configurable (134.79.46.44:5100).
Fallback: SPEC counters
read_source: spec → 'ct 0.1' refresh (else last-count values).
Requires remote control; may pulse the shutter (sauto).
Selectable per run; EPICS chosen when the CA link is up.
Notes: EPICS path = SPEC-free live plot that survives collections. Adopted mid-beamtime with the engineer's PVs.

## Slide 10
REDUCTION & CORRECTION
2D detector images → 1D I(q)
PyFAI azimuthal integration
poni calibration + detector masks; SAXS & WAXS.
Corrections
Transmission, normalization (bstop/i0/absolute), polarization.
Metadata
CSV or PDI; operator/provenance captured per file.
Auto-reduction
Folder monitor picks up new .raw and reduces in sequence.
Notes: Existing physics-audited core. For autonomy the auto-reduction monitor is the trigger — start it per run.

## Slide 11
AVERAGING & SUBTRACTION
Pairing sample and background
Averaging
Repeated scans grouped by recipe_id + role and averaged (log-space aware).
Background subtraction
Pair by recipe_id
sample ↔ background matched on the shared id; nearest-index fallback.
Scale
Manual or auto high-q (drive residual → 0), with a QC verdict.
Auto monitor
Watches Averaged, writes Subtracted, runs unattended.
Notes: Pairing by recipe_id makes autonomous subtraction robust; auto-subtraction monitor is the unattended trigger.

## Slide 12
ML-READY OUTPUT
Fixed q-grid + a unit fix that mattered
Truncate + rebin
Every subtracted profile → fixed grid the ML model needs.
Default 0.03–0.6 Å⁻¹, 549 points, linear.
App inputs (manual + automated tabs) for range / #points / spacing / unit.
The unit fix ! -> ok
Subtracted files are Å⁻¹, but the analyzer/optimizer work in nm⁻¹.
Left unfixed → sizes 10× off, optimizer chases a phantom target (silent).
Fix: analyzer detects the header unit and converts Å⁻¹→nm⁻¹ before fitting.
Notes: A silent correctness bug the readiness audit surfaced; fixed with a unit-aware analyzer + test.

## Slide 13
ANALYSIS & OPTIMIZATION
The brain of the loop
Nanoparticle fit
Polydisperse-sphere form factor (Schulz / log-normal).
Guinier size, Porod invariant, mesophase indexing.
Returns size / PDI / phase + a 0–1 confidence.
Bayesian optimizer
numpy-only GP + Expected Improvement, Sobol cold-start.
Loss = (size−target)² + PDI penalty; confidence-weighted.
Writes next recipe → Conditions/ folder the reactor watches.
LLM is advisory only (narration / QC notes) — never in the control decision.
Notes: Analyzer = fit + optimizer + writes conditions. LLM advisory, not in the control path.

## Slide 14
PUMP FLOW CALIBRATION
Water-calibrated sensors → true ODE flow
Sensors are water-calibrated and we bypass the Dolomite FCC, so the app corrects flow itself — plus per-pump flow-health safety.
cal × factor
Linear, live-editable in the app (Pump-limits card).
Flowrate table
Power-law fit (paws-style), config — wins over cal ×.
Flow-OK check
Flags a pump whose true flow strays from setpoint.
Volume limit
Ends the run if a pump over-delivers.
Notes: No serial command sets liquid type; FCC selection doesn't carry to our app. We correct in-app (factor or power-law table, matching paws) + flow-OK & volume-limit safety.

## Slide 15
THIS BEAMTIME
What we did
1
Beamline bring-up
Validated live reads, temperature set/read, shutter open/close, and a real 2D collection — via standalone safe test tools.
2
EPICS live reads
Added SPEC-independent caget monitoring (temp/i0/bstop) that survives collections.
3
Hardware workarounds
Flush via ode_dilution (no-air ode_flush skipped); per-pump flow calibration; low-flow limits.
4
Pipeline for ML
q-truncation to 0.03–0.6 Å⁻¹ / 549 pts; unit-aware analyzer fix; data_dir auto-follows the hub.
5
Simplify & harden
Removed ramp arming; app inputs authoritative; Python-3.9 startup fix for two apps.
6
Audits & readiness
Full autonomous-loop audit, safety audit, and an operator runbook.
Notes: The accomplishments slide. Six themes; each a concrete shipped change with tests.

## Slide 16
QUALITY
Tested and audited
135
automated tests passing
Coverage
Golden end-to-end pipeline regression (raw→reduce→average→subtract→analysis).
Reactor safety: E-stop pumps-only, hard SPEC guard, collection-vs-command order.
Beamline driver (mock + SPEC), EPICS read, macro streaming, collect timing.
Flow calibration (factor + power-law), flow-OK, volume limit, pump skip.
Optimizer, nanoparticle fit, background pairing, ML truncation & units.
Audited safe
No persistent config / motor / file changes to the beamline — verified against the manuals & code.
Notes: The suite guards physics (golden regression) and safety behaviors, so refactors don't silently break the loop.

## Slide 17
DOCUMENTATION
Handover material produced
Autonomous run readiness
end-to-end audit + go/no-go
Autonomous run steps
operator runbook
Beamline safety audit
what we send / never send to SPEC
Reactor map
symptom → file/function troubleshooting
Reactor setup + hardware
software + fluidics/beamline wiring
Pump flow calibration
cal × factor & power-law table how-to
Beamline testing runbook
reads/temp/shutter/collect tools
Autopilot design
proposed pipeline orchestrator (next)
Notes: Everything written down: setup, safety, troubleshooting, calibration, and the go/no-go checklist.

## Slide 18
OPEN ITEMS
What's left before hands-off autonomy
Autonomous collection trigger
Shutter & collect work via the test tools; firing during a live reactor run is the current debug (backend/timing/thread path).
In progress
Quality gate before optimizer
Analyzer fits every subtracted file; add a confidence/QC gate so bad subtractions can't bias the campaign.
Recommended
Run-end: measurement vs duration
A file.averaged event can end a run before the set duration; add a toggle to run the full duration when wanted.
Decision
Manual monitor starts
Reduction/viewer/background monitors + campaign start by hand — the autopilot orchestrator removes this.
Planned
Notes: Honest status: loop wired and each stage validated; firing collection inside a live run is the open item, plus hardening/UX items with plans.

## Slide 19
NEXT STEPS
Toward reliable, hands-off runs
1
Finish the collection trigger
Resolve the run-time fire path; verify a full cycle end-to-end on the rig.
2
Autopilot orchestrator
One process runs reduce→average→subtract in sequence, event-driven off the reactor; pre-flight validation + unified log.
3
Confidence gate + safe defaults
Only feed high-confidence fits to the optimizer; make app duration authoritative when desired.
4
Environment & polish
Move the control PC to Python 3.12; finalize EPICS as default; official SLAC template.
Notes: Priorities: finish trigger, then orchestrator (removes manual steps), then confidence gate, then env/polish. Orchestrator design doc already written.

## Slide 20
Summary
A tested, documented closed-loop
synthesis + SAXS platform — beamline-ready.
Bench-validated beamline control · autonomous data path wired · 135 tests green.
Remaining: fire collection inside a live run, then the autopilot orchestrator.
Thanks to the BL1-5 staff and the synthesis team.
SLAC National Accelerator Laboratory · SSRL BL1-5
Notes: Close: ready and validated, short path to fully hands-off autonomy. Thank the beamline staff.